## BOM
# | Subsystem | Item | Function / Notes | Make or Buy | Suggested Source / Standard | Package Reference | Qty (fill in) | Unit Cost (fill in) | Line Total
1 | 1. Lyophilisation Vessel & Chamber | Downward-conical vessel shell, double-jacketed | Product-contact pressure/vacuum boundary; holds product through freeze+dry | Contract out (specialist fabricator) for vacuum-rated unit; Make in-house for atmospheric mockup only | ASME BPVC Sec. VIII Div.1 UG-28 to UG-30 / ASME BPE / PED (EU) | 05, 09, 10
2 | 1. Lyophilisation Vessel & Chamber | Top lid & clamp ring | Closes vessel top; carries instrument/nozzle ports | Make | Sanitary tri-clamp style closure | 05
3 | 1. Lyophilisation Vessel & Chamber | Vacuum-rated sight glass | Visual inspection during cycle | Buy | Vacuum-glass manufacturer, rated to full vacuum + cryo temp | 05, 09
4 | 1. Lyophilisation Vessel & Chamber | Lid instrument/nozzle bosses (illumination, vent valve, spray nozzle mounts) | Mounting points for lid-mounted fittings | Make | Per NL2026893B1 lid fitting layout | 05
5 | 1. Lyophilisation Vessel & Chamber | Vessel support ring / lugs | Structural support to frame | Make | Standard vessel-support engineering practice | 05, 08
6 | 2. Agitator & Drive | Agitator shaft & mixing element (screw/paddle/ribbon) | Wall-scraping agitation; 0.5-15 mm (pref. 1-10 mm) wall clearance | Make | Per NL1022668C2 disclosed clearance range | 05, 10
7 | 2. Agitator & Drive | Magnetic coupling drive assembly | Seal-less torque transmission through vessel wall | Buy | Mag-drive coupling manufacturer (pressure-reactor/pharma-mixer market) | 05, 10
8 | 2. Agitator & Drive | Rotary shaft seal (alternative to mag coupling) | Lower-cost sealed shaft penetration | Buy | Vacuum/cryo-rated mechanical seal | 05, 10
9 | 2. Agitator & Drive | Agitator drive motor + gearbox + VFD | Variable-speed rotation of agitator | Buy | Standard industrial motor/VFD | 05, 08, 11
10 | 2. Agitator & Drive | External bearing housing | Keeps drive bearings outside cold zone | Make | Per NL1022668C2 | 05
11 | 2. Agitator & Drive | Agitator torque/current monitoring sensor | In-process freeze/load indicator | Buy | Standard drive instrumentation | 08, 11
12 | 3. Vessel Jacket & TCU | TCU (Temperature Control Unit) package | Heats/cools jacket heat-transfer fluid across full cycle | Buy (or assemble from purchased components) | Industrial process-chiller/TCU manufacturer | 06, 10
13 | 3. Vessel Jacket & TCU | Heat-transfer fluid (silicone oil) | Wide-range (-55C to +50C+) jacket fluid | Buy (consumable) | Silicone heat-transfer oil, freeze-dryer/reactor grade | 06
14 | 3. Vessel Jacket & TCU | Jacket circulation pump | Circulates heat-transfer fluid through jacket | Buy | Standard process pump, fluid/viscosity rated | 06
15 | 3. Vessel Jacket & TCU | Jacket supply/return piping + insulation | Delivers fluid to/from jacket | Make (piping) / Buy (insulation) | Insulated per ambient-loss and safety practice | 07, 08
16 | 4. Refrigeration System | Compressor(s) - cascade low & high stage | Refrigeration duty for TCU (and condenser if shared plant) | Buy | Cascade refrigeration, e.g. ethane/R404A/ammonia low stage | 06, 09
17 | 4. Refrigeration System | Refrigerant charge | Working fluid for refrigeration circuit | Buy (licensed handling) | Selection per GWP/flammability/toxicity tradeoffs | 06, 09
18 | 4. Refrigeration System | Refrigeration-side heat exchangers (condensers/evaporators) | Refrigerant-to-fluid heat exchange | Buy | Standard refrigeration equipment | 06
19 | 4. Refrigeration System | Refrigerant piping, receivers, safety devices | Refrigerant containment and relief | Contract (licensed refrigeration technician) | ASHRAE 15 / IIAR / local refrigeration code | 09
20 | 4. Refrigeration System | Oxygen-deficiency monitor | Required if LN2/dry ice used for direct freezing | Buy | Confined-space cryogen safety practice | 09
21 | 5. Vacuum System | Dry screw pump / Roots blower train (2+ stage) | Primary vacuum generation, oil-free | Buy | Per NL1022668C2: 2+ stage system suited to ~50 L charge | 06, 10, 11
22 | 5. Vacuum System | Backing pump | Final stage of pump train | Buy | Matched to primary pump | 06
23 | 5. Vacuum System | Pirani gauge (chamber + condenser, x2) | Thermal-conductivity pressure gauge; end-of-drying detection | Buy | Comparative pressure measurement practice (AAPS PharmSciTech 2017) | 06, 08
24 | 5. Vacuum System | Capacitance manometer (chamber + condenser, x2) | True absolute pressure; primary control signal | Buy | Same as above | 06, 08
25 | 5. Vacuum System | Vacuum isolation valve, vessel-to-collector (open/closed/throttle) | Controls vapor path per NL2026893B1 recirculation strategy | Buy | Bellows/diaphragm-sealed vacuum valve | 06, 08
26 | 5. Vacuum System | Bypass valve & piping around isolation valve | Routes dust-laden vapor to filter once valve closed | Buy/Make | Per NL2026893B1 | 06
27 | 5. Vacuum System | Vent/break valve + sterile air or N2 filter | Controlled return to atmosphere | Buy | Aseptic vacuum-break practice | 06
28 | 5. Vacuum System | Vacuum piping (short runs, large diameter) | Connects vessel, collector, condenser, pump | Make/Buy | Standard vacuum-engineering conductance practice | 08
29 | 6. Material Collector / Dust Filter | Filter housing | External vacuum-line housing per NL2026893B1 | Make | Per NL2026893B1 geometry description | 06, 10
30 | 6. Material Collector / Dust Filter | Filter screen / cartridge element | Traps dry product dust from vapor stream | Buy (consumable) | Mesh size selected per product particle size | 06, 10
31 | 6. Material Collector / Dust Filter | Filter blowback / purge valve + gas supply | Dislodges caked powder from filter screen | Buy | Per NL2026893B1 | 06
32 | 6. Material Collector / Dust Filter | Heated collection hopper / recipient, double-jacketed | Post-collection secondary drying of collected powder | Make | Per NL2026893B1 | 06
33 | 6. Material Collector / Dust Filter | Filter differential-pressure transmitter | Detects filter loading / triggers blowback | Buy | Standard filter-monitoring instrumentation | 08
34 | 7. Bottom Discharge Valve | Ball-segment discharge valve | Full-emptying bottom discharge | Buy | Hosokawa ISEM-type or equivalent sanitary valve (GEMU, ITT Pure-Flo, Bardiani, etc.) | 05, 10
35 | 7. Bottom Discharge Valve | Pneumatic valve actuator | Automated open/close of discharge valve | Buy | Standard sanitary valve actuator | 08
36 | 7. Bottom Discharge Valve | Product collection canister | Receives discharged dry powder | Buy/Make | Per AFD architecture | 04
37 | 8. CIP System | CIP skid (pump, chemical dosing, heater) | Automated wash-cycle utility supply | Buy | EHEDG-consistent CIP design practice | 07, 10
38 | 8. CIP System | Rotational spray nozzle(s) in vessel | Cleans internal vessel/lid surfaces | Buy | Per NL1022668C2 'rotational washing nozzle' | 07, 10
39 | 8. CIP System | CIP supply/return piping | Connects CIP skid to vessel | Make/Buy | BPE sanitary tube practice | 07
40 | 9. SIP System (optional) | Clean steam generator or plant clean-steam interface | Supplies >121C saturated steam for sterilization | Buy | Per Hosokawa CIP/SIP technology page | 07
41 | 9. SIP System (optional) | Steam distribution piping & traps | Delivers steam through vessel/jacket/lines | Buy | Standard clean-steam distribution practice | 07
42 | 9. SIP System (optional) | SIP-rated RTDs (multiple distribution points) | Heat-distribution/penetration validation | Buy | SIP validation requirement | 08, 14
43 | 10. Seals, Gaskets & Insulation | Lid/door gasket | Vacuum + (optional) SIP-rated seal | Buy | EPDM (steam-resistant) or PTFE-faced silicone | 07
44 | 10. Seals, Gaskets & Insulation | Port/nozzle seals (sight glass, vent, spray nozzle, etc.) | Independent leak points at every lid penetration | Buy | Per port-specific manufacturer rating | 07
45 | 10. Seals, Gaskets & Insulation | Valve stem seals (bellows/diaphragm type) | Vacuum-integrity on all vacuum-side valves | Buy | Vacuum-engineering best practice | 06, 07
46 | 10. Seals, Gaskets & Insulation | Closed-cell foam / mineral wool insulation | Thermal insulation of vessel/jacket/piping | Buy | Standard industrial insulation practice | 07
47 | 10. Seals, Gaskets & Insulation | Removable stainless cladding | Weatherproof, cleanable, inspectable insulation cover | Make | Standard sanitary-plant cladding practice | 07
48 | 11. Instrumentation (general) | Jacket RTDs, supply & return (PT100) | TCU control loop feedback | Buy | Standard process instrumentation | 08
49 | 11. Instrumentation (general) | Product temperature probe(s) | In-process product temperature monitoring | Buy | Placement is a design/prototyping problem, not off-the-shelf | 08
50 | 11. Instrumentation (general) | Level / fill sensor (or sight glass for prototype) | Batch fill monitoring | Buy | Capacitance or load-cell based for automation | 08
51 | 12. Valves & Piping (general) | Sanitary diaphragm valves (CIP/utility isolation) | Cleanable process/utility isolation | Buy | Standard sanitary diaphragm valve practice | 08
52 | 12. Valves & Piping (general) | BPE-grade sanitary tube & fittings | Product-contact and CIP/SIP piping | Buy | ASME BPE tube/fitting standard | 05, 08
53 | 12. Valves & Piping (general) | Orbital welding (equipment or contracted service) | Crevice-free sanitary tube joining | Contract or Buy (equipment) if investing in-house | ASME BPE joining practice | 08, 10
54 | 13. Control System (PLC/HMI/SCADA) | PLC | Recipe sequencing, interlocks, I/O | Buy | e.g., Siemens S7-1200/1500 or Allen-Bradley CompactLogix class | 08
55 | 13. Control System (PLC/HMI/SCADA) | HMI panel | Operator interface, trending, alarms | Buy | Standard industrial HMI | 08
56 | 13. Control System (PLC/HMI/SCADA) | Data historian / batch-record software | Continuous batch data logging | Buy | 21 CFR Part 11 considerations if pursuing GMP path | 08, 14
57 | 13. Control System (PLC/HMI/SCADA) | I/O modules, panel wiring, enclosure | Physical control panel build | Buy (components) / Make (assembly) | Standard industrial panel-building practice | 08
58 | 14. Electrical System | Motor VFDs | Variable-speed drive for agitator/pumps | Buy | Standard industrial VFD | 08
59 | 14. Electrical System | Control power supplies (24VDC) | Logic-level power distribution | Buy | Standard panel-building practice | 08
60 | 14. Electrical System | ATEX-rated motors/junction boxes/lighting (dust zone) | Ignition-source control in powder-handling zone | Buy | ATEX Zone 20/21/22 equipment category per Dust Hazard Analysis | 08, 09
61 | 14. Electrical System | E-stop circuit hardware (hardwired, PLC-independent) | Safety-critical emergency stop | Buy | Standard machine-safety practice | 09
62 | 14. Electrical System | Main disconnect & panel protection | Electrical safety isolation | Buy | Standard electrical code practice | 08
63 | 15. Structural Frame | Skid/frame (structural steel) | Supports vessel, drive, and services | Make | Standard structural fabrication practice | 08, 10
64 | 15. Structural Frame | Access platform / steps (if vessel tall) | Safe operator access to top fittings | Make | Standard industrial access practice | 08
65 | 15. Structural Frame | Vibration isolation mounts | Isolates rotating equipment vibration from vessel/frame | Buy | Protects agitator wall-clearance tolerance and instrument readings | 08
66 | 16. Utility Management Skid / Interface | Utility junction panel/manifold | Single documented interface to building services | Make | Per Hosokawa's named 'Utility Management Skid' concept | 07
67 | 16. Utility Management Skid / Interface | Compressed air/N2 supply connection & filtration | Instrument air, blow-nozzle gas, vacuum-break gas | Buy | Standard plant-utility practice | 07
68 | 16. Utility Management Skid / Interface | Process/purified water connection | CIP rinse water supply | Buy | Standard plant-utility practice | 07
69 | 16. Utility Management Skid / Interface | Drain connection | CIP/SIP effluent and condensate removal | Make | Standard plant-utility practice | 07
70 | 17. Safety Systems | Vacuum relief/vent path (manual + interlocked) | Safe return to atmosphere before any access | Make/Buy | Mandatory before opening any vacuum vessel | 09
71 | 17. Safety Systems | Lid/access interlock switch | Prevents opening under vacuum | Buy | Standard machine-safety interlock practice | 09
72 | 17. Safety Systems | Dust Hazard Analysis (DHA) - professional service | Documented combustible-dust risk assessment (not a physical part) | Contract (qualified process-safety engineer) | NFPA 652 / ATEX DSEAR requirement | 09
73 | 17. Safety Systems | Explosion venting or suppression hardware (per DHA outcome) | Physical explosion protection at material collector/discharge zone | Buy (specialist) | NFPA 68/69 or equivalent, sized per product Kst/Pmax | 09
74 | 17. Safety Systems | Combustible-gas / O2-deficiency monitors | Refrigerant leak / cryogen asphyxiation detection | Buy | Standard refrigeration/cryogen safety practice | 09
Legend:
Yellow cells = fill in for your own build (Qty, Unit Cost). Left blank intentionally; no prices or quantities are invented in this package.
'Package Reference' = the numbered file in this documentation set that explains the rationale/citation for that line.

## Summary
AFD-Style Freeze Dryer — Replica Build BOM Summary
# | Subsystem | Item Count | Make Count | Buy Count | Contract/Other Count
1 | 1. Lyophilisation Vessel & Chamber | 5 | 3 | 1 | 1
2 | 2. Agitator & Drive | 6 | 2 | 4 | 0
3 | 3. Vessel Jacket & TCU | 4 | 1 | 3 | 0
4 | 4. Refrigeration System | 5 | 0 | 4 | 1
5 | 5. Vacuum System | 8 | 1 | 7 | 0
6 | 6. Material Collector / Dust Filter | 5 | 2 | 3 | 0
7 | 7. Bottom Discharge Valve | 3 | 0 | 3 | 0
8 | 8. CIP System | 3 | 1 | 2 | 0
9 | 9. SIP System (optional) | 3 | 0 | 3 | 0
10 | 10. Seals, Gaskets & Insulation | 5 | 1 | 4 | 0
11 | 11. Instrumentation (general) | 3 | 0 | 3 | 0
12 | 12. Valves & Piping (general) | 3 | 0 | 2 | 1
13 | 13. Control System (PLC/HMI/SCADA) | 4 | 0 | 4 | 0
14 | 14. Electrical System | 5 | 0 | 5 | 0
15 | 15. Structural Frame | 3 | 2 | 1 | 0
16 | 16. Utility Management Skid / Interface | 4 | 2 | 2 | 0
17 | 17. Safety Systems | 5 | 1 | 3 | 1
| TOTAL | 74 | 16 | 54 | 4
Estimated Total Build Cost (fill in BOM sheet's Qty/Unit Cost columns first):
0